## Sheet1
Ballpark figures for equipment cost for a basic COVID-19 diagnostic lab (for RNA extraction, basic RNA QC, qRT-PCR, sample storage, and sterilization/decontamination ONLY)
DOES NOT INCLUDE: kits, relevant supplies/glasswares, PPE, additional sample/RNA quality control equipment, additional equipment for tests such as conventional PCR, and any form of laboratory renovation required to conform to BSL-2 standards
Equipment list is based on a standard qRT-PCR research lab. List is subject to updates based on information on COVID-19 testing protocols
Throughput: at least 180 samples/day, depending on the protocol
Item # | Equipment Description | Purpose | Quantity | Unit Cost | Subtotal | Company/Brands available locally
1 | Autoclave (40-60L)** | Sterilization of reagents/plasticware and decontamination of clinical waste | 1 | 350000 | 350000 | Nuve Daihan Scientific
2 | Micropipette Set 1x 1-10ul 1x 10-100ul 1x 100-1000ul | Liquid handling | 5 | 40000 | 200000 | Vitlab Witeg Nichiryo Themo-Fisher *
3 | Class IIA2 Biosafety Cabinet** | Safe handling of clinical samples | 1 | 450000 | 450000 | Thermo-Fisher Labconco* Esco*
4 | Digital weighing Scale 0.01g readability** | Weighing of reagents and balancing of centrifuge loads | 1 | 170000 | 170000 | Adam Equipment Sartorius Labconco*
5 | Glass Door Refrigerator at least 400L** | Storage of reagents/kits | 1 | 150000 | 150000 | Biobase Evermed* Esco* Thermo-Fisher*
6 | Drying Oven** | Drying of cleaned/sterilized plasticware and glassware | 1 | 300000 | 300000 | Memmert Shellab Nuve IKA
7 | Microcentrifuge ** 24 placer for 1.5ml tubes at least 13,000g max speed | For sample preparation and RNA isolation | 1 | 250000 | 250000 | Thermo-Fisher Orto-Alresa Sartorius*
8 | Chill racks/Cooling racks: 1.5 ml tubes | Keeps samples, RNA, and PCR reagents stable in cold conditions | 4 | 8000 | 32000 | Eppendorf Cleaver Scientific Isofreeze Tarsons
9 | Chill racks/Cooling racks: 0.2ml tubes (plate compatible) | Keeps prepared PCR reactions in stable cold conditions | 4 | 8000 | 32000 | Eppendorf Cleaver Scientific Isofreeze Tarsons
10 | Mini-microcentrifuge for 1.5ml tubes | Pulse spinning of reagents/samples (for preventing contamination) | 1 | 70000 | 70000 | Bio-rad Cleaver Scientific
11 | Mini-microcentrifuge for 0.2ml tubes OR microplate spinner | Pulse spinning of reagents/samples (for preventing contamination) | 1 | 80000 | 80000 | Bio-rad Cleaver Scientific
12 | Nanospectrophotometer** | Quality control of extracted RNA & troubleshooting of qPCR reactions | 1 | 950000 | 950000 | Thermo-Fisher
13 | Real-time PCR thermal cycler w/ laptop*** | Runs COVID-19 test PLEASE CHECK COMPATIBILITY WITH TEST KIT | 1 | 2500000 | 2500000 | Thermo-Fisher Bio-Rad
14 | Frost-free -20°C Freezer at least 400L** | Long term storage of samples, RNA, and reagents | 2 | 280000 | 560000 | Evermed Thermo-Fisher*
15 | Vortex mixer | Mixing of reagents and samples | 1 | 60000 | 60000 | Vortex Genie Thermo Fisher* Labnet*
16 | Water distiller (Portable or Cabinet-type) | Source of clean water for some reagents, for washing of glasswares | 1 | 150000 | 150000 | Generic Thermo-Fisher*
| *note: may be more expensive than indicated budget. Check with your local distributor |  |  | TOTAL | 6304000
**sensitive equipment - ensure that unit comes with AVR
***critical testing equipment - ensure that equipment comes with AVR AND UPS